## Slide 1: Guide to using the Velos Pro Mass Spectrometer
OHSU Proteomics Shared Resource
Keith Zientek
April 2020
Notes: Welcome. This brief tutorial describes OHSU researchers using the Thermo Velos mass spectrometer in the OHSU PSR core.

## Slide 2: Velos Pro Mass Spectrometer
The Velos Pro is a user-operated linear ion trap mass spectrometer managed by the OHSU Proteomics Shared Resource Facility
The Velos can be operated by any OHSU researcher after training by PSR personnel
Notes: This instrument in the core is a basic linear ion trap from Thermo Finnegan. It is quite robust and is available for use by any OHSU researcher who has completed training from the OHSU PSR core staff.

## Slide 3: Velos Pro Mass Spectrometer
Available Experiments include:
  Mass determination of small peptides (<5000 daltons)
  HPLC separation using C18 and C4 column chemistries
  Measurement of fragment ions in MS/MS experiments
  Whole mass measurements of larger polypeptides (5k-50k daltons
Notes: The OHSU PSR core has high-resolution instruments for shotgun identification and quantification proteomics experiments. The linear trap is available for basic mass spectrometry measurements to support research such as expression and purification of protein constructs. Some of the more common uses are peptide (less than 5000 Da) mass measurements and smaller protein whole masses (5K to 50K Da). A variety of dissociation methods are available to fragment and sequence peptides, too. Reverse-phase liquid chromatography using C18 (peptides) or C4 (proteins) columns are available for purification/separations.

## Slide 4: Direct Infusion
The fastest method of analysis involves direct infusion of a sample into the mass spectrometer.
The Velos Pro is equipped with an integrated infusion pump capable of 1-20 uL/min.
Typical sample volumes are 50-200 uL in a compatible buffer system for mass spectrometry
Notes: If you have sufficiently purified starting material, HPLC may not be needed. Direct infusion via syringe is the fastest way to get data on your samples. The samples need to be in solutions that are compatible with the highly sensitive instrument. Sample volumes and concentrations also need to be within specified ranges.

## Slide 5: Buffer Selection for Direct Infusion
Care must be taken to provide the sample with compatible buffers for mass spectrometry analysis.
  Sample must be suspended in volatile acidic buffers to facilitate electrospray ionization.
  Examples of incompatible buffers: Detergents such as Tween or SDS, glycine buffer, involatile salts such as NaCl and phsophate, PEG, DMF, glycerol and insoluble material such as particulates or cellular debris.
  A typical solvent for direct infusion is 50%/50% methanol/0.2% formic acid in water.
  Samples must be filtered through a 0.2 um filter to remove particulates.
  The PSR core has compatible solvents available for use and can provide advice on sample preparation for direct infusion.
Notes: Here are some of the buffer constraints. The main concerns are that particulates and detergents, although common in regular sample handling, are NOT compatible with the instrument. The buffers also need some source of hydrogen ions (H+ in acidic solutions) to create ions that the mass spectrometer manipulates and measure.

## Slide 6: Column cleanup and HPLC analysis
Samples can undergo a cleanup process prior to analysis if this is needed.
PSR can provide a C18 or C4 reverse phase column for separation or desalting prior to mass spectrometry analysis.
Consult with PSR Personnel for additional information.
Notes: HPLC chromatography with reverse phase C18 or C4 columns are available for separations of analytes of interest. It may not be practical to start with very highly purified samples, and HPLC may be used to produce cleaner analytes for detailed measurements.

## Slide 7: Reserving Time on the Velos Pro Instrument
The Velos Pro is typically operated in a first come first served basis.
The instrument can be reserved by contacting PSR personnel at least 24 hours in advance.
Notes: After you are trained and qualified to operate the Velos Pro instrument, it is available on a first come first served basis. The instrument can also be reserved for planned experiments.

## Slide 8: Data Analysis
Training in the data analysis software will be performed by PSR personnel.
Users are expected to perform their own data analysis with advice and support provided if needed.
The PSR has a computer center in Room MRBA528 available for data analysis.
Notes: Operation of the instrument requires training by the OHSU PSR core staff. Analysis of the generated data is also part of the instrument use. The OHSU PSR core has proprietary vendor software on core computers available for use in room MRBA528. The PSR core staff provide training in the software use. The analysis of the types of data generated by the Velos Pro is within the capabilities of trained users, and users are expected to perform their own analysis. The core staff is happy to answer any questions and provide help.